STANDARD OPERATION PROCEDURE

Faculty of Biosciences, NMBU

Method name: Total phosphorus

BIOVIT-nr: Arb1168

1. Introduction

In acidic solution, phosphate ions bind to ammonium molybdate and form a

Phospho-molybdate complex which is measured spectrophotometrically at 340 nm.

The absorbance is proportional to the concentration of phosphorus in the sample.

2. Reagents

- Phosphorus (inorganic) assay (PHOS PH 8328)

- Clinical Chemistry Calibrator Level 2 CAL2350

- Clinical Chemistry Calibrator Level 3 CAL2351

- Quality control (Assayed Chemistry Premium Plus Level 2, Cat NoHN1530)

- Quality control (Assayed Chemistry Premium Plus Level 3, Cat NoHE1532).

  - Enough for 350 tests. Store in a refrigerator.

- Control test:  chicken feed, Felleskjøpet

- 6 M HCl and 1 M HCl

3. Risk assessment

The reagent contains sulfuric acid that can irritate.

In case of spillage: Rinse with plenty of water and seek medical help.

HCl is corrosive, wear gloves, lab coat and work in a fume hood.

4. Equipment

- Filter (Merck Eurolab 108282-70), Kebo (if meal samples)

- Scintillating glass (20mL) and lid for boiling

- Automatic pipettes

- Centrifuge

- Muffle furnace (550 °C)

- 100 mL measuring flask

- Eppendorf tube without lid- 2 mL

- RX Daytona+ instrument

5. Special remarks

- It is especially important to use phosphate-free equipment!

  - All glass equipment is washed with 1M HCl and rinsed with MilliQ water

- Limit of detection: 0,091 mmol/L

- Limit of quantification: 0,22 mol/L

- Linearity: 0,22 – 129 mmol/L

- Repeatability (within-day Precision): 1,6-2,6 %

- Reproducibility (between-day Precision): 3,1-4,6 %

6.  Sample material

Degree of grinding 0.5 mm.

7. Work procedure

- Weigh in about 500 mg sample in a scintillation glass

- Place the glasses in the muffle oven at 550° C for 16 hours

- Take out the samples and let them cool down

- Add 5 mL 6M HCl to each sample

- Boil the samples on a heating-plate for 20 minutes.

  - Use a glass lid to avoid loss of sample

- Transfer the sample to a 100 mL measuring flask (pre-washed in 1M HCl)

- Rinse the scintillation glass with 7.5 mL MilliQ water and add to the measuring flask

- Mix well and leave the samples overnight in the acid/water mixture.

- Next day: Dilute the sample in the measuring flask up to 100 mL with MilliQ water

- Let the samples stand until they are clear (Some samples must be centrifuged)

- The sample is analyzed at RX Daytona+ with kit PH8328 (Randox)

  - See separate manual for use of the instrument

8. Calculation

Total amount of phosphorus in g/kg sample:

Read value mmol/L * 30,97 mg/mmol * 0.1 L = mg P/g sample = g/kg

Weighed sample g

Mw phosphorus = 30.97 g/mol

Dilution - 0.1 L